## Sheet1
Comparison Traditional Barrel connector vs Spring-Pin connector
Topic | element | Spring-Pin Connector (McNamara-Davis) |  | Barrel Connector (HP-Dell style)
|  | Cable connector | Device Connector | Cable connector | Device connector
Physical Size | width footprint | ~5mm width x 22 mm outside of lock | 4.318 mm x 16.21 mm depression + latch | 7mm outer tip diameter. 11 mm diameter for Dell beyond barrel. | ~11 mm minimum width plus mounting material
| length footprint | ~ 22.86mm to outside of locking ring | ~ 22.86mm to outside of locking ring | 7mm outer tip diameter. 11 mm diameter for Dell beyond barrel. | ~11 mm minimum width plus mounting material
| Depth | 14.31mm in current form | 1.1mm + pcb. | 45mm for Dell | >12.5mm insert+ connector backing for >15mm.
Current capacity |  | 9A Each power pin is rated at 9A each contact. Assume redundent pins rated at half power. | 9 A rated target pin with concave surface. Dual pins for best contact points. | 7-10A? | 7-10A?
Life Cycle expectations |  | 1,000,000 spring pin cycles - Latch pin cycles to be determined | Contact with convex plates. | 500 - 5000 cycles | 500 - 5000 cycles
Signal pin |  | one or two small center pins | one or two small center pins | small center pin | small center pin
Communications |  | Single or Dual wires would allow differential signaling. | Single or Dual wires would allow differential signaling. | single wire with respect to -power | Single center pin.
Light retention mechanism for trip resistance |  | Light connection at low lock tension will pull out in any direction through have dome angle | Retention is through spring of locking clamps. | Not capable other than at normal to connection surface | No relieve for off axis pull
High retention capability: for hard to access locations. Must be able to be retained to X newtons. (Cat Proof) |  | High retention force by lowering the lock ring lower to surface to clamp the clams. Even in clamped position, cable must release prior to damage to device. |  | Capable of retension of straight pull. High retention off axis.
insulation creepage and clearance |  | 2mm to meet IEC 60950-1 Type 3 polution and 60601-1 basic insulation. | 2mm creepage and clearance | Appears to be ~2.6mm
Goal- Not mate with existing designs |  | Will not mate with existing desgns. | Will not mate with existing design. No confusion. | Would mate with many existing design with different voltages | Many existing power supplies of all voltages can plug in.
Power capability for 10-130W |  | Yes | Yes | Yes | Yes
Capable of sourcing and sinking power |  | Yes | Yes | Yes | Yes
Consider future smaller profile devices |  | Could support 5mm thick device | small thin footprint for small devices. | Minimum footprint 11mm + material
Water resistant connection seal to keep water and coffee, tea,rain, adult beverages out |  | Built into connector | basically sealed interface. Intrface cleaned with tissue or sanitising washes. | Not available but could possibly be made available | Difficult. Easily filled with water or dirt. Difficult to remove blockage
User appeal |  | High appeal as shown by Apple | High appeal - sealed small indent + latch indent. | Same old Design
Relative Cost |  | Moderate Cost |  | Low Cost
Testing - Power Cord/Cable
Device to Device | Target
Power connection - each phase - includes 2 connectors. Measured at device side of connection | <100 milliohms - after 1000 cycles
Cable Capacitance - power connections - Positive connection to Negative connection completed cable | < 1nf
Cable Inductance - power leads - per connection | < 1millihenry
Single pin signal connection - cable resistance | <100 milliohms after 1000 cycles
Single pin signal connection - impedence to negative power condutor | 45-60 ohms
Two signal Pin differential impedence | 90 - 120 ohms
Signal insertion loss | 6db
EMI | FCC B or better
EMC | ISO --
Test current | 15 Amps
Contact Resistance | <20 milliohms
Communication Requirements
Testing - Mechanical Release Cable pull | Release Force - Newtons
For springpin - X axis is long axis of connector, Y axis is inplane perpendicular to X axis, Z axis is mutually perpendicular out of plane. Symmetry is assume for +/- Y axis and +/- X axis. For Barrel connector Z axis is in line with barrel.
Z axis pull 0 X, 0 Y component - Low retention position
Z axis pull 0 X, 10 degree Y component - Low retention position
Z axis pull 0 X, 30 degree Y component - Low retention position
Z axis pull 0 X, 45 degree Y component - Low retention position
Z axis pull 0 X, 60 degree Y component - Low retention position
Z axis pull 0 X, 90 degree Y component - Low retention position
Z axis pull 10 degree X, 0 Y component - Low retention position
Z axis pull 30 degree X, 0 Y component - Low retention position
Z axis pull 45 degree X, 0 Y component - Low retention position
Z axis pull 60 degree X, 0 Y component - Low retention position
Z axis pull 90 degree X, 0 Y component - Low retention position
Z axis pull 0 X, 0 Y component - High retention position
Z axis pull 0 X, 10 degree Y component - High retention position
Z axis pull 0 X, 30 degree Y component - High retention position
Z axis pull 0 X, 45 degree Y component - High retention position
Z axis pull 0 X, 60 degree Y component - High retention position
Z axis pull 0 X, 90 degree Y component - High retention position
Z axis pull 10 degree X, 0 Y component - High retention position
Z axis pull 30 degree X, 0 Y component - High retention position
Z axis pull 45 degree X, 0 Y component - High retention position
Z axis pull 60 degree X, 0 Y component - High retention position
Z axis pull 90 degree X, 0 Y component - High retention position